## Slide 1
THE METHOD
RESULTS
CONCLUSION
APPLICATIONS
FUTURE RESEARCH
The central aim of the project is to use a method called Frequency Modulation to increase the metrology of the Shear Force Acoustic Near-Field Microscopy (SANM).
SANM consists of a probe made up of a fabricated tip and a tuning fork that vibrates at a frequency yielding maximum amplitude. As the probe penetrates the liquid, the amplitude decreases. Unfortunately, the amplitude decreases due to two other factors:
Tiny environmental fluctuations caused by temperature and humidity
The frequency yielding maximum amplitude shifts in another medium (air to liquid)
Frequency Modulation Feedback Control for Near-Field Acoustic Characterization of Mesoscopic Fluid Films
By Pramith Devulapalli
BACKGROUND INFORMATION
The fundamentals of the Scanning Probe Microscopy (SPM) and the long standing scientific challenges provided by the mesoscopic fluid constitute the breadth of the project.
Fig. 1 represents a basic diagram of the Atomic Force Microscopy
Fig. 2 represents a basic diagram of the Scanning Tunneling Microscopy
Fig. 3 represents a basic diagram of the Near-Field Scanning Optical Microscopy
The three pictures above represent the three major types of SPM. In the AFM, a fabricated tip is attached to a cantilever to map the surface. As the tip scans along the topography of the surface, the reflected laser fluctuates and these are recorded by the photodiode. Similarly, the STM consists of a fabricated tip attached vertically to a tube and measures the current between the sample and tip. The fluctuations of the current describe the topography of the sample. In NSOM, a fabricated optical fiber coated in a certain metal sits vertically. A beam of light is excited through the tip and the reflections of the light are captured to study the sample.
Fig. 4 represents the properties and a depiction of a liquid confined between two solid objects, which is also known as the mesoscopic fluid.
The thin film of water squeezed between the ball and the surface comprises of the mesoscopic fluid. The dynamics of the liquid are remarkably different from those in the macroscopic form. The liquid’s viscosity or resistance against an applied force greatly increases, the relaxation time is prolonged when a disturbance is applied and the liquid takes time to return to equilibrium, and systematic molecular ordering appears.
OBJECTIVE
Image created by me
Solid
substrate
Confined mesoscopic fluid
Acoustic
waves
Acoustic sensor
~ 20 nm
Lateral
oscillations
Acoustic SANM signal
Probe’s amplitude of oscillation
Driver and detection
of mechanical motion
Probe
Fig. 5 represents the model of the SANM instrumentation. The probe is a fabricated tip attached to a tuning fork excited at a frequency to engender vibrations. These vibrations disturb the mesoscopic fluid and produces effects that are recorded.
Pramith Devulapalli
The STM current provides the tip-sample distance to accurately measure SANM’s approach to the substrate.
Two
Methods
The power transistor is a method to implement temperature control to record the dynamics of the fluid at various conditions.
Fig. 6 represents the integration of SANM and frequency modulation
Chart: SR-850 Lock-In Amplifier Amplitude vs. Frequency
Chart: SR 850 Lock-In Amplifier Phase vs. Frequency
Chart: Digital Lock-In Amplifier Phase vs. Time
Chart: Digital Lock-In Amplifier Amplitude vs. Time
Chart: Digital Lock-In Amplifier Frequency vs. Time
Chart: Frequency Modulation PID Error Output vs. Time
Chart: Frequency Modulation VCO Output Frequency vs. Time
Nanoelectromechanical systems, NEMS, are on the verge of transforming technology. One of the main obstacles in using NEMS is the high frictional force when these tiny machines work. This problem has hindered further development and through studies of mesoscopic fluids, nano-lubricants can be manufactured to counteract friction on the nanoscale.
In conclusion, frequency modulation with SANM is a much better instrumentation that can potentially characterize the mesoscopic fluid. Each individual part of the phase locked-loop has reliable functionality and the cooperation of these parts displays unity. By comparison to the SR-850 Lock-In Amplifier, the digital Lock-In Amplifier is a low cost and effective piece of instrumentation.
The folding and unfolding of proteins has huge implications in learning about certain neurodegenerative disorders such as Parkinson's and Alzheimer's. Hydrophobic amino acids drive protein folding and observing the effects of two hydrophobic surfaces interact with an apparatus such as SANM could potentially lead to solutions.
Apparatus such as SANM can map and diagnose the topography of a cell on the resolution of nanometers and angstroms. Cell topography and three dimensional images of a cell membrane can reveal many key aspects of cell functions. The probe can contain chemical substances to stimulate cells.
This set of graphs represent the reliable and accurate functionality of the Lock-In Amplifier. All the data was taken in a frequency sweep ranging from 32.740kHz-32.790kHz and the frequency versus time graph displays the sweep. In this range, the resonance frequency of the tuning fork lies at 140 seconds at the frequency of 32.756 kHz. The Digital Lock-In graphs display a 90 degree phase shift and a maximum amplitude of 58 nanoamps at 140 seconds. The corresponding graphs taken by the SR-850 commercially verified Lock-In Amplifier suggest very similar data results about the same tuning fork. The detected phase, amplitude, and frequency are respectively 90 degrees, 60 nanoamps, and a frequency between 32.755kHz-32.758kHz.
The two graphs above and below display the proper functionality of the Voltage Controlled Oscillator and PID microcontroller. The graph above represents the PID Error Output versus time. The main point of interest in the PID error graph is the plateau located around 140 seconds where the PID error stabilizes at 85.3 units. In the graphs of the Lock-In, the resonance frequency was located at approximately 140 seconds. At that point, the phase shift is 90 degrees and meets the set point of the PID and consequently, the PID outputs a zero error which shows the significance of the plateau. The VCO graph responds to the PID error in an orderly fashion and displays the plateau.
Lock-In Amplifier
Quartz Tuning Fork
PID Microcontroller
Voltage Controlled Oscillator
Frequency Modulation/Phase Locked Loop
The PID measures a variable, the phase identified by the Lock-In, and compares it to a set point or a goal. The set point is the ideal condition the PID searches for and in this case, a ninety degree phase shift (refer to stage 4). Depending on the difference between the measured variable and the set point, the PID outputs a calculated error that runs the operation of the next stage, the VCO. The calculated error depends on integral and derivative operations.
The Lock-In Amplifier is the heart of frequency modulation. The job of the Lock-In Amplifier is to extract phase, amplitude, and frequency from any electrical signal. The Lock-In performs its job by obtaining two signals, a reference signal and an input signal, and applies a series of mathematical operations to decode the key components.
The quartz tuning fork is the center of SANM with frequency modulation. A key principle of quartz tuning forks is that at a certain frequency, the resonance frequency, a ninety degree phase shift and maximum vibrations are the result. By keeping the ninety degree phase shift constant between the input and output of the tuning fork, resonance is guaranteed and the decrease in amplitude is attributed to the mesoscopic layer.
The Voltage Controlled Oscillator is the third stage of frequency modulation. The VCO receives the PID output error as an input signal. Depending on the PID output error, the VCO adds or subtracts a certain number of Hertz to the nominal frequency. The nominal frequency is the base frequency where the VCO starts. With the VCO relying on the PID error to shift the frequency upwards or downwards, the maximum amplitude of the tuning fork is guaranteed.
Stage 1
Stage 2
Stage 3
Stage 4
SR-850 Lock-In
Digital Lock-In
http://teachers.stanford.edu/activities/SPMReference/SPMReference.pdf
http://www.iap.tuwien.ac.at/www/surface/stm_gallery/stm_schematic
http://pharmrev.aspetjournals.org/content/60/1/43/F10.expansion
http://physics.aps.org/articles/v3/73
Image created by me
Image created by me
Image created by me
http://ansom.research.pdx.edu/introduction/sfm/
Image created by me
http://phys.org/news10452.html
http://chemwiki.ucdavis.edu/Biological_Chemistry/Proteins/Protein_Structure/Protein_Folding
http://www.biologyjunction.com/cell_model_instructions.htm
http://www.ieap.uni-kiel.de/surface/ag-kipp/stm/stm.htm
http://www.zmescience.com/other/science-abc/science-abc-the-transistor-15092010/